HOSPITAL:

Respiratory Care Services Policy and Procedure Manual

| Policy and Procedure: | Universal Bite BlockTM |
| --- | --- |
| Area: Respiratory Care Services | Performed by: Respiratory Care Practitioners |
| Policy Number: | Approved by: |
|  | Date: |
| Current Effective Date | Approved by: |
|  | Date: |
| Review Date | Approved by: |
|  | Date: |
| Revised Date | Approved by: |
|  | Date: |

POLICY

This policy assures the standardized of use of the Universal Bite BlockTM for use as positive protection from biting of the endotracheal tube and the pilot balloon tube.

PURPOSE

The purpose is to provide an easily implemented protocol to be used by the Respiratory Care Practitioner with effective guidelines and consistent instruction for use and patient application of the Universal Bite BlockTM with an endotracheal tube or similar device.

DEFINITION

The Universal Bite Block™ fits around an endotracheal tube and prevents biting of endotracheal tubes or similar devices and helps provide protection of the pilot balloon tube on the endotracheal tube.

B&B Universal Bite Block/Adult - PN#11160 fits endotracheal tube sizes 6.0 - 9.5 mm

B&B Universal Bite Block/Pediatric - PN#11170 fits endotracheal tube sizes 4.0 - 6.0 mm

SETTINGS

The Universal Bite Block™ should be placed and secured to the endotracheal tube in an environment in which the patient can be physiologically monitored and in which emergency equipment and appropriately trained health care providers with airway management skills are immediately available.

EQUIPMENT

Universal Bite BlockTM (includes the Universal Bite BlockTM, Releasable Cable Tie and Blue Cap), Alcohol Swabs, Bandage Scissors, Skin and Oral Care Supplies, Suction setup with an appropriate size suction catheter, an oral suction attachment. Use of B&B endotracheal tube holders are recommended with the Universal Bite Block.

PROCEDURE

A: Application Procedure/Preparation

1. 	Remove the Universal Bite Block from the package and release the Cable Tie and 	remove the Blue Safety Cap (Figure 1).

2.	To release the Cable Tie (Figure 2), push down the small tab on the Cable Tie hub and 	push the excess tab through the hub.

3.	Secure the endotracheal tube per instructions or hospital policy.

4.	After securing the endotracheal tube, gently pull the pilot tube in a straight line toward 	the end of the endotracheal tube.

B: Application Procedure/Action Steps

1.	Below the tape on the endotracheal tube, start with the Universal Bite Block at a slight 	angle & push the back of the Universal Bite Block onto the endotracheal tube.

2.	Keeping the Universal Bite Block at a slight angle gently slide it onto the tube.

3.  	Make sure the pilot balloon tube line is between the Universal Bite Block and the 	endotracheal tube.

3.	Keep the pilot balloon line taut; gently pull the Universal Bite Block forward up the 	endotracheal tube until the retaining flange is over the taped portion of the endotracheal 	tube.

4.	At this point rotate the retaining flange from a 6 O’clock position to a 3 O’clock position 	so that it is lateral to the endotracheal tube.

5.	Take the elongated tab of the Cable Tie & pass it around the endotracheal tube tape & 	through its Locking Hub.

6.	Firmly pull the long leg of the Cable Tie, while securely holding the Cable Tie Hub. Pull 	until secure.

7.	If using the B&B StabilTubeTM or B&B LockTiteTM, place the cable tie around the 	endotracheal tube above the first cable tie.

8.	Upon ascertaining correct placement of the Universal Bite Block, the cable tie may be 	trimmed to within 1/4 inch of the Hub on 11160 and flush on 11170.

9.	Replace the Blue Safety Cap on the end of the trimmed releasable Cable Tie.

10. 	To adjust the position of the Universal Bite Block, release the Cable Tie and adjust the 	Universal Bite Block/E.T. Tube position & retighten the Cable Tie. To release the Cable 	Tie (Figure 2), push down the small tab on the Cable Tie hub and 	push the excess tab 	through the hub.

C: Documentation

Chart the Time, Date, endotracheal tube size and cm marking at the appropriate 	anatomical landmarks on the Ventilator flow Sheet.

D: Removal and Changing of Universal Bite Block

1. 	Prior to changing the position or removing the Universal Cable Tie, release the Cable Tie 	as shown in Figure 2.

2. 	While stabilizing the endotracheal tube, gently move the position of the Universal Bite 	Block. Secure the Cable Tie and replace the Blue Cap.

E: Precautions and Adverse Events

1.	CAUTION: This should be used with caution on patients with capped teeth, bridges or 	false teeth, as permanent damage may occur to teeth. Benefits should be weighed 	against possible damage to the teeth by the ordering physician.

2.	Possible Adverse effects include but are not limited to lip sores. Try to keep the 	Universal Bite Block toward the center of mouth.

3.	Periodically inspect the Universal Bite Block and patient every shift, prn or more 	frequently in case of diaphoresis or copious oral secretions.  Make sure the Universal  	Bite Block is in the proper position and working properly.

4.	Should cuff pressure filling problems occur, gently pull the pilot balloon line taut to 	remove any kinks.

5.	Pediatric endotracheal tubes can distort if over tightened. Check the endotracheal tube 	after tightening the Cable-Tie and if distorted, remove cable tie with nail clippers, and 	tape to endotracheal tube.

6. 	Do not pull the Releasable Cable Tie tight with any other device except two fingers. If 	pulled too tight the endotracheal tube may hourglass and alter the inner diameter of the 	endotracheal 	tube. This can be noted by a significant dip in the endotracheal tube 	surface on either side of the Cable Tie.

7. 	CAUTION: If the Cable Tie is not used, the Tab must be securely taped to the 	endotracheal tube.

8.	The Universal Bite Block should be changed every three days for hygienic 	purposes.

9.	The Universal Bite Block is intended for single patient use.

Figure 1

Figure 2

REFERENCES

1.	AARC Clinical Practice Guideline/Management of Airway Emergencies

Respir Care 1995; 40(7):749-760)

# Text boxes

Blue Safety Cap

Cable Release Hub

Cable Tie