HOSPITAL:

Respiratory Care Services Policy and Procedure Manual

| Policy and Procedure Title: | Bubble CPAP for Neonates |
| --- | --- |
| Area: Respiratory Care Services | Performed by: Respiratory Care Practitioners |
| Policy Number: | Approved by: |
|  | Date: |
| Current Effective Date | Approved by: |
|  | Date: |
| Review Date | Approved by: |
|  | Date: |
| Revised Date | Approved by: |
|  | Date: |

POLICY

This policy assures the standardized use of The B&B Bubbler Water Seal CPAP Valve that provides a proven method of delivering 0 – 10 cm H2O continuous positive airway pressure (CPAP) recommended for use with infants weighing < 10 kg.

PURPOSE

The purpose of Bubble CPAP is to deliver an optimal CPAP level for the prevention or alveolar collapse, the recruitment of collapsed alveoli, improved ventilation/perfusion ratios, and to provide adequate arterial oxygenation.

DEFINITION

The B&B Bubbler is intended to provide CPAP to spontaneously breathing neonates and infants who require breathing support due to conditions associated with prematurity or other conditions where CPAP is prescribed by a physician.

SPECIFICATIONS AND OPERATION

The B&B Bubbler has an internal overflow chamber which limits the water level to the targeted and desired height throughout the course of therapy. Eternal ports enable the user to drain water in the overflow chamber and add water to the system without disengaging the device, allowing the B&B Bubbler to remain a closed system.

The Bubbler is a double-walled vessel. The inner vessel controls the CPAP level. There is a small passage which opens to the outer vessel where the oscillations of the bubbling of water is dampened. The mean water level can be accurately read in the outer vessel without interrupting the delivery of CPAP to the patient.

SPECIFICATIONS

- Connector: ISO 5356-1:2004 compliant 15mm x 22mm, 360 degrees single swivel 	connector

- Pressure: Range 0-10 cmH2O

- Accuracy: +/- 1.0 cmH2O

- Gas Flow Range: 1 to 12 LPM

- Material: Polypropylene and ABS

- Operating Temperature: 50 degrees Fahrenheit to 104 degrees Fahrenheit (10 	degrees Celsius to 40 degrees Celsius)

- Storage Temperature: -40 degrees Fahrenheit to 140 degrees Fahrenheit (-40 degrees 	Celsius to 60 degrees Celsius)

SETTINGS

To ensure patient safety the device is intended to be used under medical supervision by trained healthcare professionals.

INDICATIONS

- Provide CPAP, in a hospital setting, to spontaneously breathing neonates and infants            < 10kg.

2.	Infants spontaneously breathing, in respiratory distress, requiring oxygen.

3.	Respiratory Distress Syndrome

4. 	Transient Tachypnea of the Newborn (TTN)

5.	Respiratory support after extubation.

CAUTIONS

The B&B Bubbler is intended to be used by qualified medical personnel trained in oxygen delivery, ventilation, cardiopulmonary monitoring, and assessment of the newborn.

Evaporation or condensation of water may occur during operation of this device. Frequent monitoring of the water level is required to assure consistent delivery of set pressure.

CONTRAINDICATIONS

1.	Contraindicated in patients not requiring elevated lung pressure therapy

2.	Choanal atresia

3.	Nasopharyngeal obstruction

4.	Cleft palate

WARNINGS

CPAP therapy may have an adverse effect on the cardiopulmonary status of the patient and must be prescribed by a licensed physician.

Use of this device at gas flow rates greater than 12 LPM is not recommended and may result in the delivery of higher pressures than what is set.

EQUIPMENT

1. 	Bubble CPAP setup; The B&B Bubbler (Bubble generator), heated wire circuit                                    	(recommended), Pressure Relief Manifold (pop-off valve), humidifier, and sterile water.

2.	CPAP prongs, size appropriate

3.	CPAP headgear, size appropriate

4.	Medical air/oxygen blender, medical oxygen and medical air, flow meters, quick 	connects, and mounting (B&B Pole Clamp) equipment

5.	Skin protectant barrier

6.	Pressure manometer

7.	Suction equipment

PROCEDURE

1. 	Verify physician order

2.	Size the patient for the appropriate size NCPAP prongs and headgear. Prongs should fit snug in the infant’s nares, but do not place undue pressure or cause blanching on the septum.

3.	Nasal CPAP setup:

NCPAP with B&B Bubbler Water Seal CPAP Valve:

- Inspect the device for damage and discard if damaged.

- Fill the water reservoir to the solid fill line with sterile water or sterile         0.25% acetic acid solution (Refer to F and J). Add water to Bubbler          through either the gas inlet adapter (H) or by instilling water through         the water level fill port (B).

- Slide the Bubbler into a 30mm pole or a rail-mount bracket to assure        that the device remains in an upright position using the equipment            mount located on the rear side of the Bubbler (I).

- While holding the CPAP dial (C), press the Setting Lock Button (D)          and rotate dial counterclockwise to align the red arrow with the                 prescribed CPAP level (G and A). Release the Setting Lock Button to       lock the Dial in place.

- Assemble the patient circuit and set flow between 5 to 8 LPM,                  assuring that there is no resistance in the patient circuit with the               smaller prong sizes. Set the PEEP level +5 cmH2O (unless                 otherwise instructed) and connect expiratory limb of the circuit to the        gas inlet (H). Connect an airway pressure manometer to the circuit           between the patient interface and the inspiratory limb of the circuit.           Connect a pressure relief manifold (pop-off valve) between the                 blended gas supply and the inlet port of a humidifier.

- Prior to patient application, occlude circuit to assess for leaks in the system and assure adequate bubbling. Set FiO2 prescribed by physician.

- Suction the nares of the patient to assure patency.

- With the infant in the supine position, head elevated about 30                   degrees and neck supported with a small roll; position the prongs             gently into the nares so that there is no upward pressure on the                septum. Place prongs curve side down, and connect them to the              patient circuit. Adjust the angle at the corners, to curve the prongs            slightly, matching the curved plane across the upper lip. To minimize        the risk of irritation to the nasal septum, the bridge between the                prongs should not be in direct contact with the infant’s nose.

- Secure the airway in place. Assure ideal headgear size, position of tubing holders, and position of infant prongs.

- The patient circuit should be adjusted so that the weight of the circuit       does not place tension on the prongs.

- Set the humidifier temperature at 37 degrees. Frequently assess the        patient circuit for rainout and water condensation and drain the circuit       as needed.

- Document the patient setup.

- CPAP function and patient tolerance should be evaluated frequently.

- Monitor NCPAP routinely for the following:

  - Set FiO2

  - Liter flow set 5 to 8 LPM (less flow indicated with prong size 0       and size 1)

  - PEEP set at +5 cmH2O, or prescribed pressure

  - Correct water level in Bubbler (Bubble Generator)

  - Humidifier temperature stable at 37 degrees and rainout in           the circuit is drained

  - Nasal prong size and positioning are correct

  - Headgear fit snugly, but not too tight

  - Septum is intact and checked for any skin breakdown

  - Skin integrity

  - Water is continuously bubbling

  - Patient work of breathing, breath sounds, respiratory rate,             heart rate, and secretions

  - Once a shift, remove the infant’s headgear and inspect              head for evidence of breakdown. Assure that the areas are dry         and free of compromise before reapplying the headgear and       prong apparatus

4.	Recommended care with CPAP Therapy

1.	Obtain a baseline blood gas within the first hour of initiation of CPAP and

PRN with any change in clinical status including an increase in oxygen                                               			requirement, an increase in the severity, or frequency of apnea and 				bradycadia, or increase in respiratory distress.

2.         Monitor blood gas frequently per physician order.

3.         Assess chest x-rays routinely, per physician order.

4.         Assure that an orogastric tube is placed to vent air accumulation in the

abdomen.

5.          Rotate patient position per hospital protocol.